## Worksheet 1
Laboratory and sample preparation
Laboratory name | Dept of Earth Sciences, ETH Zurich
Sample type/mineral | Carbonates, mostly calcite
Sample preparation | Thin sections or chips mounted in epoxy
Imaging | CL8200 Mk5-2 optical cathodoluminescence system
Laser ablation system
Make, model and type | ASI (Resonetics) RESOlution S155
Ablation cell and volume | Laurin Technic, two-volume cell, effective volume ca. 1 cm3
Laser wavelength (nm) | 193 nm
Pulse width (ns) | 25 ns
Fluence (J cm−2) | ∼ 1.8 J cm−2
Repetition rate (Hz) | Variable, see data for actual value
Ablation duration (s) | 40 s
Ablation pit depth ∕ ablation rate | Variable, equivalent to 0.09–0.15 µm per pulse
Spot diameter (µm) nominal/actual | Variable, see data for actual value
Sampling mode/pattern | Static spot ablation
Carrier gas | 100 % He in the cell, Ar make-up gas combined in cell above ablation in funnel
Cell carrier gas flow (L min−1) | 0.5 L min−1
ICP-MS Instrument
Make, model and type | Thermo Element XR, Sector-field single collector ICP-MS, with high-capacity interface pump
Sample introduction | Ablated aerosol direct
RF power (W) | 1350–1550 W (optimised daily)
Make-up gas flow | 0.90–1.05 L min−1 Ar (optimised daily) 2 mL min−1 N2
Detection system | Triple (pulse-counting, analogue, Faraday) cross-calibrated daily with U 238, fixed analogue counting factor (ACF) value, all isotopes usually in pulse-counting only (
Masses measured (amu) | 202, 204, 206, 207, 208, 232, 235, 238
Integration time per peak/dwell times (ms) | 11 ms (all masses) except: 206, 207 (50 ms)
Total integration time per output data point (s) | 0.174 s
Sensitivity/efficiency (%, element) | ∼ 1 % U
Dead time (ns) | 25
Typical oxide rate (ThO ∕ Th) | 0.18 %
Typical doubly charged rate (Ba++/Ba+) | 3.50 %
Gas blank | 20 s
Calibration strategy | NIST614 glass standard as primary reference material for drift and Pb–Pb ratios; WC-1 carbonate standard for matrix matching of 206Pb∕238U; ASH-15-D and JT carbonate for quality control
Reference material info | NIST614 (concentration data: Jochum et al. 2011; Pb isotopes: Baker et al. 2004); WC-1 (Roberts et al., 2017); ASH-15-D (Nuriel et al., 2020); JT (characterised in this work)
Data processing package used | Iolite 2.5, VisualAge for integration, interval selection and gas blank correction only. In-house spreadsheet data processing. IsoplotR (Vermeesch 2018) for isochrons, intercept ages and initial Pb compositions
Correction for LIEF | No LIEF correction (ratio of mean intensities 207Pb∕206Pb or mean of uncorrected ratios 238U∕206Pb used)
Mass discrimination | Normalised to reference material (sample standard bracketing)
Common-Pb correction, | None applied; ages calculated from regressions in Tera–Wasserburg concordia plots.
composition and uncertainty
Uncertainty level and | Intercept ages are quoted at 2s absolute; propagation is by quadratic addition. Counting statistics
propagation | uncertainty are propagated to the 207Pb∕206Pb ratio, together with the uncertainty of the RM value and the uncertainty of repeated measurements. The uncertainty value for lower-intercept isochron ages includes uncertainties from the RM and asystematic uncertainties, estimated in this work to be 2.5 %. Decay constant uncertainties are neglected.
Quality control/validation | ASH 15D: mean of 206Pb∕238U intercept ages: 2.885±0.076 Ma (2s, MSWD = 0.99, n=30) (0.9 % weighted average uncertainty (internal), 2.0 % total external uncert.)
JT – mean of 206Pb∕238U intercept ages: 13.70±0.37 Ma (2s, MSWD = 1.03, n=16) (1.15 % weighted average uncertainty (internal), 2.5 % total external uncert.)
Systematic uncertainty for propagation is 2.5 % (2s)